## Slide 1: DDS and Laser Locking
Meir Alon
Complex Physics

## Slide 2: Outline
PLL-Phase Locked Loop
Fundamentals.
Phase Noise.
PLL Lab Examples.
DDS –Direct Digital Synthesizer
Fundamentals
Phase Trancation error
DDS Lab examples
Conclusions

## Slide 3
Simple Signal Definition:
X(t)=a(t))
X(t)=a(t)t+θ(t))
t+ θ(t) Total Phase
θ(t) Excess Phase
+ Total Freqeuncy
Excess Freqeuncy
SIGNALS - DEFINITIONS
W=
Frequency
Phase
π
t
2π
3π

## Slide 4: PLL-BLOCK DIAGRAM

## Slide 5: PLL Phase Locked Loop
Basic blocks of a PLL :
Phase
Detector
vco
REF
VCO
Vout
Low Pass Filter
Vin
Vout
Generates a voltage compared to the phase Different
Generates an RF frequency propotional to a voltage
Decay unwanted signals
And maintane stability.

## Slide 6: Loop with LPF
In this diagram we see a LPF is insereted
1)To reduce un wanted frequencys
2) To make better PLL Response.
Phase
Detector
REF
VCO
vco
Low Pass Filter
Vout
N

## Slide 7: Phase Detector
So the Phase Detector compare the Phase between the signals
Phase
Detector
REF
VCO
Vout
And we can say :

## Slide 8: PLL-ANALYSIES:Phase Detector
Assuming
(t)=
(t)=
=-
After Mixing(with filtering) we get:
(t)=*cos(-- )
When the loop is locked we assume that the is constant with time and small :
So: =constant
we get:
(t)=* cos()=(*cos()
(- )
-
v
-=0 =

## Slide 9: VCO:
=+
(t)=
=A(t))
=A(t)t+
vco
w1
w0
V1
Wout

## Slide 10: VCO Dynamics
By assuming linear approximation Between we can write:
=+
=
So the Excess phase of the VCO is:
(t)=(s)
So the transfer function of the VCO is:
(s)==

## Slide 11: A PLL
Basic Block Diagram of Phase Locked Loop:
Phase
Detector
REF
VCO
vco
We wish to align the output phase with the phase of the Reference.
At the start we have two signals that are not synchronized in phase we insert them into the PD and the PD need to output a certain voltage which will bring the VCO to the desired frequency.
We have: =+ and
So: ===
to minimize the Phase Error () must be maximized .
Vout
Wout
w1
w0
V1

## Slide 12: FIRST ORDER PLL
=1
Loop Gain=
H(s)====
it is usually Needed a High Gain for good Tracking so it forces High Bandwidth ,therefore if we want Narrow bandwidth is Incompatible with good Tracking .

## Slide 13: Second Order PLL
The vast majority of PLL are Second order PLL’S which means the degree of the denominator are 2.
Transfer Function yield:
H(s)==
In order to reach this H(s) First we will define F(s):
F(s) Is defined as F(s)=+
F(s)=- =-(+)
PI=Propotional+Integral

## Slide 14: SMALL TRANSIENTS

## Slide 15: Locked Condition: Small transients in locked conditions
Let’s assume a phase step
While the input and output waveforms expressed in:
(t)=
(t)=
=t+u(-)
The output don’t change instantaneously,the VCO initialy continues to oscillate at w1.

## Slide 16
Let’s look at the waveform of this phase step:
The input phase grow cause phase difference between input and output, which creates pulses at the output of the Phase Detector.
The frequency of the VCO has increased for this transient and it accumulated a phase:
And the output settles to
t++u(t- ))

## Slide 17: Let’s look at a Frequency jump: ∆W again at t=
There is a frequency jump at the input,the VCO initially continues to oscillate at the same frequency. The PD generates wider pulses so rises with time.
As approces
Win+∆W the pulses width decreases.
Till the frequency reaches the same value.

## Slide 18: PLL Dynanics

## Slide 19: PLL-Transfer Function
Open Loop:
=
System Transfer Function:
H(s)==

## Slide 20: Bode Diagram of the Basic PLL
We chose 10khz cross over frequency.
The loop gain magnitude should be 1 at the Cross Over Frequency (Fc):
=G(S)=
Loop Gain:
=1
F(s)=()

## Slide 21: The Total system
You can see the cross over frequency at about 10khz and the phase margin of 63 degrees

## Slide 23: PLL Designed in WIS
Specification:
1)Frequency : 1000 Ghz/ 6.8 Ghz
2)External / Internal Reference frequency.
Features:
1)Frequency Hoping by Profiles selection.
2)Frequency chirp sweep.
3)Parallel mode of opearation.
Specification:
1)Range 6.8-6.85 Ghz (Rubidium Range)
2) All DDS Features
2)External / Internal Reference frequency.
Features at High Frequency:
1)Frequency Hoping by Profiles selection.
2)Frequency chirp sweep.
3)Parallel mode of opearation.
DDS Scaling

## Slide 24: PLL Designed in WIS
Specification:
1)Phase Detector Range :
Up to 300 Mhz
2)External / Internal Reference frequency.
Features:
1)PID mechanical settings.
2)Working with beats from lasers.
3)Switches for lock/not locked.
4)Remote control of the device.

## Slide 25: PLL-Phase Noise

## Slide 26: What is actually phase Noise ?
phase Noise
“Instantaneous phase noise is the accumulation of all the phase deviation that has occurred since the oscillator first started opearation.” (F.M.Gardner)
Spectral Density:
The spectral density of noise is defined as the Fourier transform of the autocorrelation function provided the function is stationary. (units=dBc/Hz)
=
The theoretically stationary issue is usually a problem looking at the autocorrelation of some signal involed.
As an example: )

## Slide 27: What is actually phase Noise ?
Which result in non stationary process.
But this issue can be overcome by making the measurement appropriate. For example measuring in certain interval period of times.
And making transmission zero at certain places in frequency.
Normalized Spectral Density of Phase Noise :

## Slide 28: Measure Phase Noise
Spectrum Analyzer
Phase
Demodulator
Spectrum
Analyzer
Log( )
To Display
)
Mixer
Band Pass Filter
Detector
To Display
)
Log( )
VBW
Filter
Frequency
Sweep
Phase Noise Analyzer

## Slide 29: Oscillator Phase Noise
-open loop transfer function:
=
Around the frequency of oscillation (+ )
We can aproximimate H(jw) with Taylor expansion:
+
(+ ))=
Since )=+1 and typically :
Resonator
Additive
Noise

## Slide 30: Q of Oscillators
In Physics :
For an LC tank, Q is an indication of how much of the energy is lost as it is transferred from the capacitor to the inductor and vice versa.

## Slide 31: Q in Oscillators
Resonance circuits have usually bandpass transfer function. The Q is defined as the sharpness of the magnitude frequency response.
According to Barkhausen criterion of oscilation the phase shift around a stable oscilating feedback circuit must be an integer multiple of
Additive noise can interfere the stable operation by altering the phase around the loop. So the oscillator shifts it’s frequency to try bring around the loop phase back to . So the response of the phase in frequency is important.
w

## Slide 32: Oscillator Phase Noise
Spectral Density:
=
To gain more insight:A(w)exp(jɸ(w))
=(+jA) exp(jɸ)
For :
=
=
Lesson Model

## Slide 33: Oscillator Phase Noise
So from the equation :
=
Some measurement showed a phase noise spectra of:
We can derive all the contribution of the noise at the VCO ( Voltage noise , Current noise ). Which are usually specified in data sheet. And use this multiplication to get the phase noise approximation.
So if we have a source of noise it is introduced as:
Resonator
Additive
Noise

## Slide 34: Good Oscillator
Low White noise contributions from all sources.
Oscillator power should be high.
Resonator Q should be large

## Slide 35: Manufacture Phase Noise Data
Usually we get from the Manufacture a graph of VCO Phase Noise.

## Slide 36: FM Noise
]
VCO

## Slide 37: Phase noise –Propagation in a PLL or a Synthesizer
H(s)==
Low Pass
E(s)==1-H(s)=
=
PFD
vco
F(s)
1/N
REFERENCE
=H(s)+E(s)=
By superposition of noises:

## Slide 38: Spectral Density of the output of the PLL:
(𝑓)= (𝑓)+ (𝑓)

## Slide 39: Noise Floor
A relation between the time jitter to the power density of the phase noise:
If is flat ( white noise):
=8
Noise from electronic circuit like PD ,buffers and dividers:
==
=8 = 8
Referred to the Multiplier input.

## Slide 40
For the case all the noises are the same in nature:
circuit performance

## Slide 41: Phase Noise Simulations
Phase Noise Simulation in ADS Keysight:

## Slide 42: Phase Noise Measurement
Phase Noise Measurement
Phase
Demodulator
Spectrum
Analyzer
Log( )
To Display
)

## Slide 43: Total Phase Noise
Design issues:
1)In a design of PLL usually inside the bandwidth of the PLL the PD and buffer Circuits dominates the Phase Noise which determines the noise floor. outside the Bandwidth the VCO dominates the output Phase Noise.
2)In designing a PLL it is important for Phase Noise Consideration to chose the Bandwidth as the cross point where the VCO Phase noise
Hits the Noise floor Phase Noise of the PLL.

## Slide 44: Laser Locking
BDU
BDU
674 nM
Frequency
Comb
Phase Detector
+ PID
Phase Detector
+ PID
OPTICS TO LAB
1.5nM
1.56nM
Stable Laser
(Outside of the comb)
(Lee Peleg Thesis)

## Slide 45: Beat from Laser
Servo Bumps
Long Pulses/ Short Pulses

## Slide 46: DDS Direct Digital Synthsizer

## Slide 47: DDS-Basic understanding
This is a block diagram which shows a good on the basic’s of DDS.The counter increments each cycle of the clock.
Counter
ROM
DAC
LPF
Fclock
RF OUT
Fewer points of one cycle
ANGLE TO AMPLITUDE

## Slide 48: Phase Accumulator
This is a block diagram of the DDS based on the Phase Accumulator ,an amplitude to Phase transformation (Look up Table) and a Digital to Analoge Converter (DAC).
ADDER
REGISTER
ROM
y
DAC
LPF
N Bits Wide
Fclock
RF OUT
M
Word

## Slide 49: Phase Accumulator
M=1
Let’s take an example for N bits wide of the Register,
if the increment is equal to 1 ,then
the register travles from 000 to 111 ,
one complete cycle is extacted from ROM.
Each clock period the output increments the phase by 2π/8 1 cycle of the sinusoid.
M=2
The accumulator overflows after 110 and
the phase changes in 2π/4 and we get
2 cycles of the sinusoid.
ADDER
REGISTER
M
Word
ROM
x
y
DAC
LPF
N Bits Wide
Fclock
RF OUT
000
010
100
110
000
001
010
011
100
101
110
111

## Slide 50
M=3
The accumulator overflows at 110 at the first cycle,111 in the second cycle and 101 at the third cycle, and starts at 000 again.so we got 3 cycles of the sinusoid.
M=4
The accumulator overflows at 100 each cycle.so we got 4 cycles of the sinusoid.
So the period of the output signal is:
n - The number of clocks the system clock should make In order to complete one period of sinusoid.
And:
So we get :
And M is the Frequency Word ,which added each time to the accumulator.
000
011
110
001
100
111
010
000
100
100
100
100
M=3
M=4

## Slide 51: AD9910 - DDS
Modes of opearations:
Single tone
Digital Ramp Generator
Parallel Data control

## Slide 52: The Phase Accumulator Digital Phase Wheel
We divide the circle to points that each point is a phase point on a sine wave and we jump from each point to another each cycle of the System Clock.
If we have cycle time of :
The Digital Phase Wheel basically Show all the phase possibilities around one cycle of a sine wave.

## Slide 53: Phase Truncation Error
DDS use only a fraction of the bits from
the output of the Accumulator and the others are ignored (truncated),this cut introduce an error called Truncation Error.
This truncation cause periodic spurs at the RF spectrum. Because of phase error.
Example:
A 8 bit Accumulator
3 bit Truncating, so 5 significant bit of phase resolution.
Each Green Bit is:
Each Red Bit is:

## Slide 54: DDS Designed in WIS
Specification:
1)Frequency Range: DC- 400Mhz
2)External / Internal Reference frequency.
Features:
1)Frequency Hoping by Profile selection.
2)Frequency chirp sweep.
3)Parallel mode of opearation.

## Slide 55: Conclusions
Taking about PLL, Fundamentals and Phase noise.
PLL Systems at the lab
Using DDS Systems and the errors involved
PLL or DDS application Dependent.

## Slide 56: Thank You

## Slide 57: Phase Modulation Signal
We reach:

## Slide 58
Assume a clean carrier has a sinusoidal phase jitter with a peak deviation of radians at a rate
The peak amplitude of each resultant sideband at is:
C is the carrier power